{"text_format":"3","workbook":"xls","n_sheets":1,"sheets":["Sheet1"],"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":"Sheet1","index":0,"n_rows":5,"n_cols":7,"rows_kept":5,"merged":[[0,0,0,6]],"flags":["formulas_cached_only"]}
["钦州港海关手持式气体检测设备需求清单"]
["序号","品名","数量","单价预算（元）","功能概述","主要技术指标","备注"]
[1,"手持式溴甲烷检测仪",1,30000,"对溴甲烷气体进行连续检测，主要用于熏蒸气体残留检测场景。","★检测气体：溴甲烷 CH3Br\n★检测量程：0-100ppm\n检测原理：PID\n▲检测精度：±2%FS\n▲分辨率：0.1ppm\n▲响应时间：＜60s\n▲零点漂移：≤ ±2% FS /年\n▲其他功能：可使用电池供电，适用于室外作业环境；方便进行数据读取，具备测量数据存储传输功能；支持声光等报警方式等。","带“★”的为必须满足内容，否则作为无效投标处理，不予接受；\n带“▲”的为重要内容，不满足将严重影响评分。"]
[2,"手持式磷化氢检测仪",1,20000,"对磷化氢气体进行连续检测，主要用于熏蒸气体残留检测场景。","★检测气体：磷化氢 PH3\n★检测量程：0-20ppm\n检测原理：EC\n▲检测精度：±2%FS\n▲分辨率：0.1ppm\n▲响应时间：＜60s\n▲零点漂移：≤ ±2% FS /年\n▲其他功能：可使用电池供电，适用于室外作业环境，方便进行数据读取，具备测量数据存储传输功能，支持声光等报警方式等。","带“★”的为必须满足内容，否则作为无效投标处理，不予接受；\n带“▲”的为重要内容，不满足将严重影响评分。"]
[3,"手持式二合一熏蒸气体检测仪",1,50000,"对多种组分气体（溴甲烷、磷化氢）进行连续检测，主要用于熏蒸气体残留检测场景。","★检测气体：溴甲烷 CH3Br，磷化氢 PH3\n★检测量程：0-100ppm，0-20ppm\n检测原理：PID，EC\n▲检测精度：±2%FS\n▲分辨率：0.1ppm\n▲响应时间：＜60s\n▲零点漂移：≤ ±2% FS /年\n▲其他功能：可使用电池供电，适用于室外作业环境，方便进行数据读取，具备测量数据存储传输功能，支持声光等报警方式等。","带“★”的为必须满足内容，否则作为无效投标处理，不予接受；\n带“▲”的为重要内容，不满足将严重影响评分。"]